LiveGuard-EHMS

Hardware, Sensor Ingestion & Edge AI Engineering Guide
Final Year Major Project

# 1. Project Architecture & End-to-End Workflow

LiveGuard-EHMS is an Edge-AI and Federated Healthcare Monitoring System designed for real-time ECG arrhythmia detection. The system operates on a staged architecture where the Edge device captures biological signals, applies digital signal filtering, detects R-peaks using the Pan-Tompkins algorithm, and runs on-device inference using a lightweight 1D-CNN (SmallConv1DGate) model.

Key System Stages:
1. Real-time Ingestion: 360 Hz ECG sampling (AD8232 via Arduino USB / Raspberry Pi MCP3008 SPI).
2. Digital Signal Processing: 0.5–40 Hz Butterworth Bandpass + 50 Hz IIR Notch filter.
3. Beat Segmentation: Pan-Tompkins QRS detector extracting normalized 180-sample beat windows.
4. Edge Inference: PyTorch Stage 1 CNN classifying Normal vs Abnormal with < 4ms latency.
5. Telemetry Broadcasting: WebSocket streaming of raw/filtered waveforms and telemetry to the Web Dashboard.
6. Federated Transfer Learning: Local fine-tuning and periodic model weight synchronization with Flower FedAvg server.

# 2. Hardware Wiring & Pinout Specifications

## 2.1 Arduino Uno / Nano Bridge Wiring (Immediate Setup)

| AD8232 Sensor Pin | Arduino Pin | Wire Color | Function |
| --- | --- | --- | --- |
| 3.3V | 3.3V | Red | Power supply (Do NOT connect to 5V) |
| GND | GND | Black | System ground |
| OUTPUT | Pin A0 | Yellow / White | Analog ECG signal output |
| LO+ | Digital Pin 2 | Blue | Leads-Off Detection Positive |
| LO- | Digital Pin 3 | Green | Leads-Off Detection Negative |
| SDN | Unconnected | — | Shutdown pin (Not needed) |

## 2.2 Raspberry Pi 4 (1GB) + MCP3008 ADC Wiring (When Pi Arrives)

| Component Pin | Raspberry Pi 4 Pin | Function / SPI Channel |
| --- | --- | --- |
| MCP3008 VDD & VREF | Pin 1 (3.3V) | Power & ADC reference voltage |
| MCP3008 AGND & DGND | Pin 6 (GND) | Analog & digital ground |
| MCP3008 CLK | Pin 23 (GPIO 11 / SPI0_SCLK) | SPI Clock |
| MCP3008 DOUT | Pin 21 (GPIO 9 / SPI0_MISO) | Master In Slave Out |
| MCP3008 DIN | Pin 19 (GPIO 10 / SPI0_MOSI) | Master Out Slave In |
| MCP3008 CS/SHDN | Pin 24 (GPIO 8 / SPI0_CE0) | Chip Select 0 |
| MCP3008 CH0 | AD8232 OUTPUT | Analog ECG channel input |
| AD8232 3.3V & GND | Pin 17 (3.3V) & Pin 14 (GND) | Sensor power |
| AD8232 LO+ & LO- | Pin 16 (GPIO 23) & Pin 18 (GPIO 24) | Leads-off detection |

# 3. 3-Lead ECG Electrode Placement on the Body

For optimal signal-to-noise ratio (SNR) and clear R-peak detection, use Einthoven's Triangle configuration:
• RA (Right Arm / Red Electrode): Place on the right infraclavicular fossa (just below right collarbone) or right inner wrist.
• LA (Left Arm / Yellow Electrode): Place on the left infraclavicular fossa (just below left collarbone) or left inner wrist.
• RL (Right Leg / Green Electrode): Place on the lower right abdominal region. This serves as the reference ground.

Practical Tip for Clean Signal: Clean application areas with alcohol wipes to remove skin oils. Keep the subject relaxed and seated during acquisition to prevent electromyographic (EMG) muscle artifacts.

# 4. Flashing Arduino Firmware & Running the Pipeline

Step 1: Open Arduino IDE and load: edge_system/arduino_bridge/liveguard_sensor_bridge.ino
Step 2: Connect Arduino via USB, select Board (Arduino Uno/Nano) and COM Port (e.g. COM3).
Step 3: Click Upload (Right Arrow button).
Step 4: Install Python dependencies on your laptop:
         pip install pyserial websockets torch scipy numpy
Step 5: Run the LiveGuard edge engine:
         python -m edge_system.run_edge --source ARDUINO --port COM3

# 5. Operating Modes Summary

1. ARDUINO Mode (--source ARDUINO):
   Reads live sensor stream from Arduino USB COM port at 360 Hz.

2. MOCK Mode (--source MOCK):
   Simulates continuous 360 Hz patient data using synthetic P-Q-R-S-T generator or MIT-BIH recordings. Ideal for developing without physical hardware.

3. RPI_SPI Mode (--source RPI_SPI):
   Direct hardware execution on Raspberry Pi 4 using spidev and MCP3008 ADC.

# 6. Division of Work & Responsibilities

Hardware & Edge Lead (You):
• Sensor interfacing (AD8232 + MAX30102 + Arduino/MCP3008).
• Digital signal processing (Bandpass, Notch filter, Pan-Tompkins QRS detection).
• Edge inference pipeline & WebSocket telemetry server.
• Web Dashboard frontend connection.

Machine Learning & Federated Learning Lead (Mirza):
• Flower Federated Learning ServerApp & ClientApp.
• Centralized baseline benchmarking & non-IID patient partitioning.
• Stage 1 threshold calibration (Recall >= 0.95).
• Stage 2 multiclass arrhythmia classifier (AAMI N, S, V, F, Q categories).